## INSTRUCTIONS
Small Field Dosimetry Calibration — Instructions & Color Key
PROTOCOL | IAEA TRS-483 (2017) + AAPM TG-155 (2021)
DETECTORS | Farmer (ref) ¦ Standard Imaging A16 ¦ Sun Nuclear Edge ¦ PTW 60019 microDiamond
ENERGIES | 6 MV WFF (Table 24) ¦ 10 MV WFF (Table 27)
SETUP DEPTH | 10 cm in water phantom, SAD or SSD=100 cm, 10×10 cm reference field
FIELD SIZES | Method A (≥3 cm): Farmer+A16+Edge+uDiam ¦ Method B (1–3 cm): Intermediate Field Chain ¦ Method C (<1 cm): Edge+uDiam mandatory 2-detector
COLOR KEY
| User Entry Cell | Enter measurement data here
| Calculated Cell | Formula — do not overwrite
| Result Cell | Final corrected result
| Label Cell | Row / column label
| 6 MV Section | Measurements at 6 MV WFF
| 10 MV Section | Measurements at 10 MV WFF
| A16 Detector Column | Standard Imaging A16 readings
| Edge Detector Column | Sun Nuclear Edge readings
| microDiamond Column | PTW 60019 microDiamond readings
| PASS | Within acceptance criterion
| FAIL / Warning | Exceeds acceptance criterion — investigate
WORKFLOW ORDER
1. SETUP — Fill equipment IDs and calibration factors (one time)
2. ENVIRONMENT — Record T and P before each measurement session
3. BEAM_QUALITY — Measure TPR20,10 for each energy
4. REF_DOSIMETRY — TRS-398 absorbed dose at reference field (Farmer)
5. DET_CONDITIONING — Pre-irradiate detectors; check leakage and dose-rate dependence
6. FWHM — Measure all field sizes with a scanning detector; record Sclin
7. OF_METHOD_A — Output factors for Sclin ≥ 3 cm
8. OF_METHOD_B — Intermediate field chain for 1 cm ≤ Sclin < 3 cm
9. OF_METHOD_C — Small field output factors for Sclin < 1 cm
10. K_FACTORS — Transcribe k_{Q_clin,Q_msr} from TRS-483 Table 24 (6MV) / Table 27 (10MV)
11. UNCERTAINTY — Review component uncertainties
12. SUMMARY — Final output factor table; check cross-detector agreement

## SETUP
Equipment Setup & Calibration Constants
Fill ONCE before commissioning. These values are referenced by all measurement sheets.
LINAC
Model / Manufacturer
Serial Number
Installation Date
TPS Name / Version
FARMER CHAMBER (Reference)
Model | Serial Number | ND,w [Gy/C] | kQ (6 MV)* | kQ (10 MV)* | Cal. Date
* kQ from TRS-398 Table T_II based on measured TPR20,10 — enter after BEAM_QUALITY sheet.
A16 — Standard Imaging A16 microionization chamber
Serial Number | Volume [cm³] | ND,w [Gy/C] | kQ (6 MV)* | kQ (10 MV)* | Cal. Date
0.016
* If no absolute calibration: A16 is used as a relative detector; k_Qmsr from TRS-483 Table 24/27.
EDGE — Sun Nuclear Edge unshielded diode
Serial Number | Sensitive Vol. [mm³] | Pre-irrad. dose [Gy] | Reference DR [MU/min] | Conditioning Date | Notes
| 0.03 | 10 | 300
Edge is used as a relative detector only — absolute dose via chain from Farmer.
microDIAMOND — PTW 60019
Serial Number | Sensitive Vol. [mm³] | Pre-irrad. dose [Gy] | Reference DR [MU/min] | Conditioning Date | Notes
| 0.004 | 5 | 300
Pre-irradiate with ≥5 Gy before first use; repeat if unused > 4 h.
ELECTROMETER
Model | Serial Number | kel (electrometer factor) | Cal. Date
1
MEASUREMENT CONDITIONS (Reference)
Reference temperature T0 [°C] |  |  | 20
Reference pressure P0 [kPa] |  |  | 101.325
Reference depth [cm] |  |  | 10
SSD or SAD [cm] |  |  | 100
Reference MU per reading |  |  | 200

## ENVIRONMENT
Environmental Conditions — Temperature, Pressure, kTP
Record at start and end of each measurement session. kTP = ((273.15+T)/(273.15+T0)) × (P0/P) where T0=20 °C, P0=101.325 kPa (values from SETUP!D rows for T0/P0).
Session # | Date | Time | T [°C] | P [kPa] | kTP | Notes / Session description
1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20

## BEAM_QUALITY
Beam Quality — TPR20,10 per Energy
Setup: Farmer chamber at isocentre (SSD+depth=100 cm), 10×10 cm field, 200 MU. Keep SSD constant; change depth only. kTP cancels in ratio if T,P stable.
6 MV WFF
Measurement | Rep 1 | Rep 2 | Rep 3 | Mean | Result
M at d = 20 cm [nC or rel]
M at d = 10 cm [nC or rel]
TPR20,10 = M20 / M10 |  |  |  |  | → Enter kQ in SETUP sheet D/E row for Farmer
Typical range for linac: TPR20,10 = 0.62–0.78. Repeat if two readings differ by >0.3%.
10 MV WFF
Measurement | Rep 1 | Rep 2 | Rep 3 | Mean | Result
M at d = 20 cm [nC or rel]
M at d = 10 cm [nC or rel]
TPR20,10 = M20 / M10 |  |  |  |  | → Enter kQ in SETUP sheet D/E row for Farmer
Typical range for linac: TPR20,10 = 0.62–0.78. Repeat if two readings differ by >0.3%.

## REF_DOSIMETRY
Reference Dosimetry — TRS-398 (Farmer Chamber)
Field: 10×10 cm ¦ Depth: 10 cm ¦ SSD=100 cm or SAD ¦ 200 MU. Complete Part A–D for each energy.
6 MV WFF
A — Environmental Correction
Session # (→ ENVIRONMENT sheet) |  |  | T [°C] from ENV |  |  | P [kPa] from ENV |  | Copy from ENVIRONMENT sheet for this session
kTP
B — Polarity Correction kpol
Collection voltage V1 [V] | 300 |  | Reversed voltage −V1 [V] | -300
| Rep 1 | Rep 2 | Rep 3 | Mean
M(+V1) [nC]
M(−V1) [nC] (enter with sign)
kpol = (M(+V1)+¦M(−V1)¦) / (2×M(+V1)) |  |  | Should be within 1.000 ± 0.005 for Farmer in MV beams.
C — Ion Recombination ks (Two-Voltage Method, pulsed beams)
V1 [V] | 300 |  | V2 [V] (V1/V2 = 2) | 150
| Rep 1 | Rep 2 | Rep 3 | Mean
M(V1) [nC]
M(V2) [nC]
r = M(V1)/M(V2)
ks (TRS-398 Eq. T.IV, pulsed beams) |  |  | Coefficients valid for V1/V2=2. ks should be <1.010 at 300/150 V.
D — Absorbed Dose to Water Dw
| Rep 1 | Rep 2 | Rep 3 | Rep 4 | Rep 5 | Mean | SD [nC] | CV [%]
M_raw [nC] at V1, +polarity
kel (electrometer factor, from SETUP)
M_corrected = M_raw × kTP × kpol × ks × kel
ND,w [Gy/C] (from SETUP Farmer row)
kQ (from SETUP / BEAM_QUALITY)
Dw [Gy per reading] = Mcorr × ND,w × kQ
Output [cGy/MU] = Dw × 100 / MU_per_reading |  | 200 | Target: 1.000 cGy/MU at reference conditions (may differ by ±2% before output tuning).
10 MV WFF
A — Environmental Correction
Session # (→ ENVIRONMENT sheet) |  |  | T [°C] from ENV |  |  | P [kPa] from ENV |  | Copy from ENVIRONMENT sheet for this session
kTP
B — Polarity Correction kpol
Collection voltage V1 [V] | 300 |  | Reversed voltage −V1 [V] | -300
| Rep 1 | Rep 2 | Rep 3 | Mean
M(+V1) [nC]
M(−V1) [nC] (enter with sign)
kpol = (M(+V1)+¦M(−V1)¦) / (2×M(+V1)) |  |  | Should be within 1.000 ± 0.005 for Farmer in MV beams.
C — Ion Recombination ks (Two-Voltage Method, pulsed beams)
V1 [V] | 300 |  | V2 [V] (V1/V2 = 2) | 150
| Rep 1 | Rep 2 | Rep 3 | Mean
M(V1) [nC]
M(V2) [nC]
r = M(V1)/M(V2)
ks (TRS-398 Eq. T.IV, pulsed beams) |  |  | Coefficients valid for V1/V2=2. ks should be <1.010 at 300/150 V.
D — Absorbed Dose to Water Dw
| Rep 1 | Rep 2 | Rep 3 | Rep 4 | Rep 5 | Mean | SD [nC] | CV [%]
M_raw [nC] at V1, +polarity
kel (electrometer factor, from SETUP)
M_corrected = M_raw × kTP × kpol × ks × kel
ND,w [Gy/C] (from SETUP Farmer row)
kQ (from SETUP / BEAM_QUALITY)
Dw [Gy per reading] = Mcorr × ND,w × kQ
Output [cGy/MU] = Dw × 100 / MU_per_reading |  | 200 | Target: 1.000 cGy/MU at reference conditions (may differ by ±2% before output tuning).

## DET_CONDITIONING
Detector Pre-conditioning, Leakage & Dose-Rate Dependence
Complete for every detector BEFORE output factor measurements. Re-do after detector storage > 4 h.
A16 — Standard Imaging A16
A — Pre-irradiation Conditioning
Required pre-irrad. dose | 0 Gy (ion chamber — no conditioning required)
Energy / field used |  | MU delivered |  | Date / Time
Conditioning readings (stability check) | Rep 1 | Rep 2 | Rep 3 | Mean | CV [%]
Reading [nC or rel] | [5 empty] | CV should be <0.3% after sufficient conditioning.
B — Leakage Current Check
| Rep 1 | Rep 2 | Rep 3 | Mean [nC] | Normal signal [nC] | Leakage [%]
Leakage (no beam)
Acceptance: leakage < 0.5% of normal signal. Subtract from all readings if >0.1%.
C — Dose-Rate Dependence
Reference dose rate: 300 MU/min. Test at 5 levels. Acceptance: <1.0% variation.
Dose rate [MU/min] | Rep 1 | Rep 2 | Rep 3 | Mean | Norm. to ref. | Deviation [%]
100 |  |  |  |  | ""
200 |  |  |  |  | ""
300 |  |  |  |  | 1.0000 | 0.000
400
600
Flag any dose rate with ¦deviation¦ > 1% — apply dose-rate correction factor if FFF beam.
EDGE — Sun Nuclear Edge (unshielded diode)
A — Pre-irradiation Conditioning
Required pre-irrad. dose | 10 Gy minimum (see SETUP for value)
Energy / field used |  | MU delivered |  | Date / Time
Conditioning readings (stability check) | Rep 1 | Rep 2 | Rep 3 | Mean | CV [%]
Reading [nC or rel] | [5 empty] | CV should be <0.3% after sufficient conditioning.
B — Leakage Current Check
| Rep 1 | Rep 2 | Rep 3 | Mean [nC] | Normal signal [nC] | Leakage [%]
Leakage (no beam)
Acceptance: leakage < 0.5% of normal signal. Subtract from all readings if >0.1%.
C — Dose-Rate Dependence
Reference dose rate: 300 MU/min. Test at 5 levels. Acceptance: <1.0% variation.
Dose rate [MU/min] | Rep 1 | Rep 2 | Rep 3 | Mean | Norm. to ref. | Deviation [%]
100 |  |  |  |  | ""
200 |  |  |  |  | ""
300 |  |  |  |  | 1.0000 | 0.000
400
600
Flag any dose rate with ¦deviation¦ > 1% — apply dose-rate correction factor if FFF beam.
microDIAMOND — PTW 60019
A — Pre-irradiation Conditioning
Required pre-irrad. dose | 5 Gy minimum (see SETUP for value)
Energy / field used |  | MU delivered |  | Date / Time
Conditioning readings (stability check) | Rep 1 | Rep 2 | Rep 3 | Mean | CV [%]
Reading [nC or rel] | [5 empty] | CV should be <0.3% after sufficient conditioning.
B — Leakage Current Check
| Rep 1 | Rep 2 | Rep 3 | Mean [nC] | Normal signal [nC] | Leakage [%]
Leakage (no beam)
Acceptance: leakage < 0.5% of normal signal. Subtract from all readings if >0.1%.
C — Dose-Rate Dependence
Reference dose rate: 300 MU/min. Test at 5 levels. Acceptance: <1.0% variation.
Dose rate [MU/min] | Rep 1 | Rep 2 | Rep 3 | Mean | Norm. to ref. | Deviation [%]
100 |  |  |  |  | ""
200 |  |  |  |  | ""
300 |  |  |  |  | 1.0000 | 0.000
400
600
Flag any dose rate with ¦deviation¦ > 1% — apply dose-rate correction factor if FFF beam.

## FWHM
FWHM Field Size Measurements
Measure inline and crossline profiles at 10 cm depth (SAD setup). FWHM = distance between 50% dose points on each side of Dmax. Sclin = √(FWHMx × FWHMy). Sclin determines which Method (A/B/C) applies.
6 MV WFF
Coll. X [cm] | Coll. Y [cm] | FWHM X [cm] | FWHM Y [cm] | Sclin [cm] | Sclin-Geo [cm] | Penumbra 20-80% [mm] | Detector used | Method | Notes
0.5 | 0.5
0.6 | 0.6
0.8 | 0.8
1 | 1
1.2 | 1.2
1.5 | 1.5
2 | 2
2.5 | 2.5
3 | 3
4 | 4
5 | 5
6 | 6
8 | 8
10 | 10
Sclin−Geo typically 1–3 mm for small fields (FWHM > geometric). Larger deviation may indicate MLC positional error.
10 MV WFF
Coll. X [cm] | Coll. Y [cm] | FWHM X [cm] | FWHM Y [cm] | Sclin [cm] | Sclin-Geo [cm] | Penumbra 20-80% [mm] | Detector used | Method | Notes
0.5 | 0.5
0.6 | 0.6
0.8 | 0.8
1 | 1
1.2 | 1.2
1.5 | 1.5
2 | 2
2.5 | 2.5
3 | 3
4 | 4
5 | 5
6 | 6
8 | 8
10 | 10
Sclin−Geo typically 1–3 mm for small fields (FWHM > geometric). Larger deviation may indicate MLC positional error.

## OF_METHOD_A
Output Factors — Method A (Sclin ≥ 3 cm)
Interleave reference (10×10) readings before and after each clinical field. Apply kTP to all readings. kpol and ks: enter from REF_DOSIMETRY sheet (Farmer); for solid-state detectors kpol=1.0000, ks=1.0000. k = k_{Q_clin,Q_msr} from K_FACTORS sheet.
6 MV WFF
Field | Sclin [cm] | kTP | Reference (10×10) — interleaved |  |  | Mref corr | Farmer |  |  |  |  | A16 |  |  |  |  | Edge |  |  |  |  | µDiamond |  |  |  |  | Ω mean | Dev [%] | Pass?
Nom [cm] | →FWHM | env→ | Mbefore | Mafter | Mref | Mcorr | M1 | M2 | M3 | Mmean | Ω_Farm | M1 | M2 | M3 | Mmean | k_A16 | Ω_A16 | M1 | M2 | M3 | Mmean | k_Edge | Ω_Edge | M1 | M2 | M3 | Mmean | k_uDm | Ω_uDm | Mean | Max% | ≤2%?
3
4
5
6
8
10
10 MV WFF
Field | Sclin [cm] | kTP | Reference (10×10) — interleaved |  |  | Mref corr | Farmer |  |  |  |  | A16 |  |  |  |  | Edge |  |  |  |  | µDiamond |  |  |  |  | Ω mean | Dev [%] | Pass?
Nom [cm] | →FWHM | env→ | Mbefore | Mafter | Mref | Mcorr | M1 | M2 | M3 | Mmean | Ω_Farm | M1 | M2 | M3 | Mmean | k_A16 | Ω_A16 | M1 | M2 | M3 | Mmean | k_Edge | Ω_Edge | M1 | M2 | M3 | Mmean | k_uDm | Ω_uDm | Mean | Max% | ≤2%?
3
4
5
6
8
10
Ω (output factor) = (M_clin × kTP_clin / M_ref_corr) × k_{Q_clin,Q_msr}. For Farmer in clin field: k≈1.000 for Sclin≥3 cm (small effect). Max deviation = (max−min)/mean across all detectors.

## OF_METHOD_B
Output Factors — Method B (Intermediate Field Chain, 1 cm ≤ Sclin < 3 cm)
Chain: 10×10 (Farmer) → Intermediate field (Farmer+detectors) → Clinical field (detectors). Ω_total = Ω(10×10→IF) × Ω(IF→clin) where Ω = (M_clin/M_ref)×k. Recommended IF size: 3×3 or 4×4 cm (ion chamber valid).
6 MV WFF
STEP 1 — Reference Field 10×10 cm (Farmer)
| Rep 1 | Rep 2 | Rep 3 | Mean | kTP | Mcorr_ref
M_ref (Farmer, 10×10) [nC or rel]
STEP 2 — Intermediate Field (Farmer, ion chamber)
IF collimator setting [cm] | 3 |  | Sclin_IF (from FWHM sheet) |  |  | k_{IF,ref} (Farmer, from K_FACTORS)
| Rep 1 | Rep 2 | Rep 3 | Mean | kTP | Mcorr_IF | Ω(ref→IF) Farmer
M_IF (Farmer) [nC or rel]
STEP 3 — Cross-calibration of Detectors at IF
| Rep 1 | Rep 2 | Rep 3 | Mean | kTP | Mcorr_IF_det | k_{IF} (det)
M_IF (A16) [nC or rel]
M_IF (Edge) [nC or rel]
M_IF (microDiamond) [nC or rel]
k_{IF}(det) = k_{Q_clin,Q_msr} for the IF field size — from K_FACTORS sheet.
STEP 4 — Clinical Field Measurements & Chain Output Factors
Nom [cm] | Sclin | kTP | A16 mean | k_A16_clin | Ω_A16 | Edge mean | k_Edge_clin | Ω_Edge | uDiam mean | k_uDm_clin | Ω_uDm
1
1.2
1.5
2
2.5
Ω_total = Ω(10×10→IF)_Farmer × (M_clin×kTP_clin)/(M_IF_det×kTP_IF) × k_clin. Positioning tolerance: ±0.5 mm for 1–3 cm fields. Compare results from all 3 detectors; accept if within 2%.
10 MV WFF
STEP 1 — Reference Field 10×10 cm (Farmer)
| Rep 1 | Rep 2 | Rep 3 | Mean | kTP | Mcorr_ref
M_ref (Farmer, 10×10) [nC or rel]
STEP 2 — Intermediate Field (Farmer, ion chamber)
IF collimator setting [cm] | 3 |  | Sclin_IF (from FWHM sheet) |  |  | k_{IF,ref} (Farmer, from K_FACTORS)
| Rep 1 | Rep 2 | Rep 3 | Mean | kTP | Mcorr_IF | Ω(ref→IF) Farmer
M_IF (Farmer) [nC or rel]
STEP 3 — Cross-calibration of Detectors at IF
| Rep 1 | Rep 2 | Rep 3 | Mean | kTP | Mcorr_IF_det | k_{IF} (det)
M_IF (A16) [nC or rel]
M_IF (Edge) [nC or rel]
M_IF (microDiamond) [nC or rel]
k_{IF}(det) = k_{Q_clin,Q_msr} for the IF field size — from K_FACTORS sheet.
STEP 4 — Clinical Field Measurements & Chain Output Factors
Nom [cm] | Sclin | kTP | A16 mean | k_A16_clin | Ω_A16 | Edge mean | k_Edge_clin | Ω_Edge | uDiam mean | k_uDm_clin | Ω_uDm
1
1.2
1.5
2
2.5
Ω_total = Ω(10×10→IF)_Farmer × (M_clin×kTP_clin)/(M_IF_det×kTP_IF) × k_clin. Positioning tolerance: ±0.5 mm for 1–3 cm fields. Compare results from all 3 detectors; accept if within 2%.

## OF_METHOD_C
Output Factors — Method C (Sclin < 1 cm, Very Small Fields)
Mandatory: ≥ 2 detector types. Positioning tolerance: ±0.1 mm. Ω = (M_clin/M_IF_det × kTP) × k_clin, where M_IF_det is from Method B cross-calibration. A16 ion chamber is typically too large for fields < 1 cm — use Edge + microDiamond. Uncertainty: ±4–6% (k=1) for fields < 1 cm.
6 MV WFF
Method C uses the same cross-calibration from Method B (Step 3). Ensure Method B is complete first and reference the M_IF values from that sheet.
Nom [cm] | Sclin [cm] | kTP | Edge M1 | Edge M2 | Edge M3 | Edge Mmean | k_Edge_clin | Ω_Edge | uDiam M1 | uDiam M2 | uDiam M3 | uDiam Mmean
[9 empty] | k_uDm_clin | Ω_uDm | Ω_mean | Dev [%]
0.5
0.6
0.8
⚠ Ω values here are RELATIVE — they require normalisation to M_IF from OF_METHOD_B for absolute output factors. For fields < 0.5 cm: uncertainty ±5–7%; considered research-level dosimetry. Acceptance criterion for 2-detector agreement: ≤ 3% for fields < 1 cm.
10 MV WFF
Method C uses the same cross-calibration from Method B (Step 3). Ensure Method B is complete first and reference the M_IF values from that sheet.
Nom [cm] | Sclin [cm] | kTP | Edge M1 | Edge M2 | Edge M3 | Edge Mmean | k_Edge_clin | Ω_Edge | uDiam M1 | uDiam M2 | uDiam M3 | uDiam Mmean
[9 empty] | k_uDm_clin | Ω_uDm | Ω_mean | Dev [%]
0.5
0.6
0.8
⚠ Ω values here are RELATIVE — they require normalisation to M_IF from OF_METHOD_B for absolute output factors. For fields < 0.5 cm: uncertainty ±5–7%; considered research-level dosimetry. Acceptance criterion for 2-detector agreement: ≤ 3% for fields < 1 cm.

## K_FACTORS
k_{Q_clin,Q_msr} Correction Factors — TRS-483 Tables 24 & 27
Transcribe k values from TRS-483 Appendix II. Table 24 = 6 MV WFF. Table 27 = 10 MV WFF. Interpolate linearly in Sclin for unlisted field sizes. k = 1.000 at the msr field (10×10 cm) by definition.
DETECTOR SOURCES: PTW 60019 → TRS-483 Table 24/27. Edge (Sun Nuclear) → check TRS-483 App. II or TG-155 literature. A16 → use nearest small IC type in TRS-483 or manufacturer data.
6 MV WFF (TRS-483 Table 24)
Sclin [cm] | A16 (k) | Edge (k) | microDiamond (k) | Farmer (for IF) | Source / Notes
0.5 |  |  |  |  | → Table 24
0.6 |  |  |  |  | → Table 24
0.8 |  |  |  |  | → Table 24
1.0 |  |  |  |  | → Table 24
1.2 |  |  |  |  | → Table 24
1.5 |  |  |  |  | → Table 24
2.0 |  |  |  |  | → Table 24
2.5 |  |  |  |  | → Table 24
3.0 |  |  |  |  | → Table 24
4.0 |  |  |  |  | → Table 24
5.0 |  |  |  |  | → Table 24
6.0 |  |  |  |  | → Table 24
8.0 |  |  |  |  | → Table 24
10.0 | 1 | 1 | 1 | 1 | k = 1.0000 by definition at msr field (10×10 cm).
For sizes not in TRS-483: interpolate linearly. Do NOT extrapolate below smallest tabulated size.
10 MV WFF (TRS-483 Table 27)
Sclin [cm] | A16 (k) | Edge (k) | microDiamond (k) | Farmer (for IF) | Source / Notes
0.5 |  |  |  |  | → Table 27
0.6 |  |  |  |  | → Table 27
0.8 |  |  |  |  | → Table 27
1.0 |  |  |  |  | → Table 27
1.2 |  |  |  |  | → Table 27
1.5 |  |  |  |  | → Table 27
2.0 |  |  |  |  | → Table 27
2.5 |  |  |  |  | → Table 27
3.0 |  |  |  |  | → Table 27
4.0 |  |  |  |  | → Table 27
5.0 |  |  |  |  | → Table 27
6.0 |  |  |  |  | → Table 27
8.0 |  |  |  |  | → Table 27
10.0 | 1 | 1 | 1 | 1 | k = 1.0000 by definition at msr field (10×10 cm).
For sizes not in TRS-483: interpolate linearly. Do NOT extrapolate below smallest tabulated size.

## UNCERTAINTY
Uncertainty Budget (k = 1, 1σ)
Based on TRS-483 §8 and TG-155 §6. u_combined = √(Σuᵢ²). Values in %. Enter measured values in yellow cells; blue cells are calculated. Cross-reference with measurement sheets for reproducibility u.
Component | Sclin ≥ 2 cm | 1 cm ≤ Sclin < 2 cm | 0.5 cm ≤ Sclin < 1 cm | Source / Basis
u1 Reference dosimetry (Dw, TRS-398) | 0.9 | 0.9 | 0.9 | TRS-483 §8.2; Farmer + calibration lab
u2 Field size measurement (FWHM, positioning) | 0.3 | 0.5 | 1 | Positioning uncertainty grows for small fields; ±0.1 mm tolerance
u3 kTP (T, P measurement) | 0.1 | 0.1 | 0.1 | Thermometer ±0.1 °C, barometer ±0.1 kPa
u4 kpol | 0.1 | 0.1 | 0.1 | TRS-398; typically <0.5% for Farmer
u5 ks (ion recombination, Farmer) | 0.1 | 0.1 | 0.1 | Two-voltage method uncertainty
u6 k_{Q_clin,Q_msr} correction factor | 0.5 | 1 | 2 | TRS-483 uncertainty on tabulated k; increases for small fields
u7 Detector energy response / fluence perturb. | 0.3 | 0.5 | 1.5 | Residual after k correction; detector-dependent
u8 Dose-rate dependence (solid-state) | 0.2 | 0.2 | 0.5 | From DET_CONDITIONING dose-rate test
u9 Reproducibility (measured) | → | → | → | From CV column in measurement sheets (Method A/B/C)
u10 Long-term stability / drift | 0.3 | 0.3 | 0.5 | Annual re-measurement; output drift since commissioning
u_combined [%] = √(Σuᵢ²) |  |  |  | Combined standard uncertainty (k=1). Multiply by 2 for 95% coverage (k=2).
U_expanded [%] (k=2, ≈95%)
TRS-483 §8 typical expanded uncertainties (k=2): ≥2 cm: ~2%; 1–2 cm (Method B): ~3–5%; <1 cm (Method C): ~8–12%. These are total output factor uncertainties.

## SUMMARY
Final Output Factor Table — Summary
Transcribe final Ω values from Method A/B/C sheets. Green = within acceptance. Cross-detector agreement ≤ 2% (Sclin ≥ 1 cm) or ≤ 3% (< 1 cm).
6 MV WFF
Nom [cm] | Sclin [cm] | Method | Ω A16 | Ω Edge | Ω µDiam | Ω Farmer | Ω Mean | Max Dev [%] | U_exp [%] | Accept?
0.5 |  | C | [6 empty] | 10
0.6 |  | C | [6 empty] | 10
0.8 |  | C | [6 empty] | 10
1.0 |  | B | [6 empty] | 4
1.2 |  | B | [6 empty] | 4
1.5 |  | B | [6 empty] | 4
2.0 |  | B | [6 empty] | 4
2.5 |  | B | [6 empty] | 4
3.0 |  | A | [6 empty] | 2
4.0 |  | A | [6 empty] | 2
5.0 |  | A | [6 empty] | 2
6.0 |  | A | [6 empty] | 2
8.0 |  | A | [6 empty] | 2
10.0 |  | A | [6 empty] | 2
All Ω are normalised to the 10×10 cm msr field (Ω = 1.0000). Ω for 10×10 cm should be 1.0000 by definition. Uncertainty column = typical expanded (k=2) — update from UNCERTAINTY sheet.
10 MV WFF
Nom [cm] | Sclin [cm] | Method | Ω A16 | Ω Edge | Ω µDiam | Ω Farmer | Ω Mean | Max Dev [%] | U_exp [%] | Accept?
0.5 |  | C | [6 empty] | 10
0.6 |  | C | [6 empty] | 10
0.8 |  | C | [6 empty] | 10
1.0 |  | B | [6 empty] | 4
1.2 |  | B | [6 empty] | 4
1.5 |  | B | [6 empty] | 4
2.0 |  | B | [6 empty] | 4
2.5 |  | B | [6 empty] | 4
3.0 |  | A | [6 empty] | 2
4.0 |  | A | [6 empty] | 2
5.0 |  | A | [6 empty] | 2
6.0 |  | A | [6 empty] | 2
8.0 |  | A | [6 empty] | 2
10.0 |  | A | [6 empty] | 2
All Ω are normalised to the 10×10 cm msr field (Ω = 1.0000). Ω for 10×10 cm should be 1.0000 by definition. Uncertainty column = typical expanded (k=2) — update from UNCERTAINTY sheet.